R/V Ronald H. Brown METADATA - 2000

Class of Data:  Surface ocean and atmospheric carbon dioxide concentrations

Dataset Identifier:  R/V Ronald H. Brown

One File:  RHB2000

Statement of how to cite dataset:

Ron Brown website: http://www.aoml.noaa.gov/ocd/gcc/rvbrown_data2000.php

These data are made freely available to the public and the scientific community

in the belief that their wide dissemination will lead to greater understanding

and new scientific insights.  The availability of these data does not

constitute publication of the data.  We rely on the ethics and integrity of

the user to assure that AOML receives fair credit for our work.  Please send

manuscripts using this data to AOML for review before they are submitted

for publication so we can insure that the quality and limitations of the data

are accurately represented.

Measurement platform identifier:  NOAA research vessel Ronald H. Brown (R104)

Cruise Information:

The Ron Brown conducted 7 major cruises in the Pacific Ocean for a total of 9 	legs.

Project Information:

The system was operated by personnel from AOML or PMEL (Pacific Marine 	Environmental Laboratory) or by the Ron Brown’s Chief Survey Tech, Jonathan 	Shannahoff.  The work was sponsored by the Underway pCO2 on Ships project of the 	NOAA climate program.

Scientist responsible for technical quality of dataset:

Rik Wanninkhof

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Rik.Wanninkhof@noaa.gov

Contact person for this dataset:

Bob Castle

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Robert.Castle@noaa.gov

Timestamp for initial submission of dataset:  09/18/08

Timestamp for the most recent update of dataset:  09/18/08

Timestamp period the dataset refers to: 3/18/2000 – 11/13/2000

Geographic area the dataset refers to:

10 S to 60 N

170 W to 94 W

2000 Cruises:

RB200001 –	Tsunami Vents

Seattle, WA to Newport, OR

March 17, 2000 to March 22, 2000

Chief Scientist – Scott Stalin

Operator – Bob Castle

RB200002 -	Vents 2000

Newport, OR to Seattle, WA

March 24, 2000 to April 9, 2000

Chief Scientist – N/A

Operator – Jonathan Shannahoff

RB200003 -	Haruphone/Tsunami

Seattle, WA to Seattle, WA

May 22, 2000 to June 4, 2000

Chief Scientist – Matt Fowler, Lt. Keith Roberts

Operator – Jonathan Shannahoff

RB200004 -	Vents Leg 1

Seattle, WA to Newport, OR

June 17, 2000 to June 27, 2000

Chief Scientist – Robert Embley

Operator – Jonathan Shannahoff

RB200005 -	Vents Leg 2

Newport, OR to Seattle, WA

June 29, 2000 to July 18, 2000

Chief Scientist – Robert Embley

Operator – Jonathan Shannahoff

RB200006 -	Vents Leg 3

Victoria, BC, Canada to Seattle, WA

July 20, 2000 to August 4, 2000

Chief Scientist – Ed Baker

Operator – Jonathan Shannahoff

RB200007 -	FOCI 2000

Dutch Harbor, AK to Dutch Harbor, AK

September 16, 2000 to September 23, 2000

Chief Scientist – Ed Cokelet

Operator – Jonathan Shannahoff

RB200008 -	NOPP 2000

Dutch Harbor, AK to San Diego, CA

September 27, 2000 to October 9, 2000

Chief Scientist – Bill Parker

Operator – Jonathan Shannahoff

RB200009 -	TAO 2000

San Diego, CA to Panama Canal

October 16, 2000 to November 13, 2000

Chief Scientist – N/A

Operator – Jonathan Shannahoff

List of variables included in this dataset:

COLUMN	HEADER			EXPLANATION

1.		GROUP/SHIP:			AOML_Brown for all underway data from the Ron Brown.

2.		CRUISE_DESIGNATION:	Cruise ID (e.g., RBYYYYnn where RB = Ron Brown, YYYY

= the four digit year, and nn = the cruise number for

that year).

3.		JD_GMT:			Decimal year day.

4.		DATE_DDMMYYYY:		GMT date.  The date format has been changed to comply

with the IOCCP recommendations.

5.		TIME_HH:MM:SS:		GMT time.

6.		LAT_DEC_DEGREE:		Latitude in decimal degrees (negative values are in 							the southern hemisphere).

7.		LONG_DEC_DEGREE:		Longitude in decimal degrees (negative values are in 							the western hemisphere).

8.		xCO2W_PPM:			Mole fraction of CO2 (dry) in the equilibrator

headspace at equilibrator temperature (Teq) in parts

per million.

9.		xCO2A_PPM:			Mole fraction of CO2 in air in parts per million.

10.		EqTEMP_C:			Temperature in equilibrator water in degrees

centigade.  Temperature in equilibrator measured with

a calibrated thermistor.

11.		PRES_EQUIL_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).

12.		SST(TSG)_C:			Temperature from the ship's thermosalinograph in

degrees centigrade.

13.		SAL(TSG)_PERMIL:		Salinity from the ship's thermosalinograph on the

Practical Salinity Scale.

14.		fCO2w,eq:			Fugacity of CO2 in the equilibrator in

microatmospheres calculated as outlined below.

15.		fCO2W@SST_uatm:		Fugacity of CO2 in sea water in microatmospheres

calculated as outlined below.

16.		fCO2A_uATM:			Fugacity of CO2 in air in microatmospheres

calculated as outlined below.

17.		dfCO2_uatm:			Sea water fCO2 - air fCO2 in microatmospheres.  This

uses the average air value for the current hour.

The following fields have been QC'ed by the CO2 group:

GROUP/SHIP

CRUISE_DESIGNATION

JD_GMT

DATE_DDMMYYYY

TIME_HH:MM:SS

LAT_DEC_DEGREE

LONG_DEC_DEGREE

xCO2W_PPM

xCO2A_PPM

EqTEMP_C

PRES_EQUIL_hPa

fCO2w,eq

fCO2W@SST_uatm

fCO2A_uATM

dfCO2_uatm

The following fields are from the ship's onboard systems and the quality of this

data cannot be verified:

SST(TSG)_C

Sal(TSG)_Permil

Narrative description of system design:

CO2 ANALYTICAL SYSTEM:

The concentration of carbon dioxide (CO2) in surface ocean water is determined

by measuring the concentration of CO2 in gas that is in contact with the water.

Surface water is pumped ~ 100 m through 7/8" Teflon tubing from an inlet

in the ship's bow to the equilibration chamber.  Water comes from the bow

intake ~4.2 m below the water line and the TSG is located close to the inlet.

When the SST is below about 20 oC, friction in the pipes and from the pump cause

heating and the Teq is higher than SST.  When the SST is higher than about 25 oC,

the ship’s air conditioning cools the water and the Teq is lower than SST.

The equilibration chamber has an enclosed volume of gas, or headspace, and a pool

of seawater that continuously overflows to a drain.  As the water flows through the

chamber, the dissolved gases (like CO2) partition between the water and the

headspace.  At equilibrium, the ratio of CO2 in the water and in the headspace is

influenced most by temperature, and that relationship is known.  By measuring

the concentration of CO2 in the headspace and the temperature in the chamber,

the partial pressure (or fugacity) of CO2 in the surface water can be calculated.

INSTRUMENT DESCRIPTION

The general principle of instrumental design can be found in Wanninkhof and Thoning

(1993), Ho et al. (1995), and Feely et al. (1999).  The concentration of CO2 in the

headspace gas is measured using the adsorption of infrared (IR) radiation, which

results from changes in the rotational and vibrational energy state of the CO2

molecule.  The LI-COR detector passes IR radiation through two 6" cells.  The

reference cell is flushed with a gas of known CO2 concentration.  The sample cell

is flushed with the headspace gas.  A vacuum-sealed, heated filament is the

broadband IR source.  The IR radiation alternates between the two cells via a

chopping shutter disc. An optical filter selects an adsorption band specific

for CO2 (4.26 micron) to reach the detector.  The solid state (lead selenide)

detector is kept at -12 degrees °C for excellent stability and low signal

noise (less than 0.2 ppm).

Several steps are taken to reduce interferences and to increase the accuracy

of the measurements.  After the equilibration chamber, the headspace travels

through a drying trap to remove water vapor.  During each analysis, the

headspace gas is compared to a reference gas of known concentration.  To

improve the accuracy of the measurements, three different gaseous standards

for CO2 are analyzed once an hour instead of the headspace gas.

Analyzer:  LI-COR 6251 (analog output) infrared (IR) analyzer.

Method of Analysis:  Differential analyses relative to the low standard.  Measures

dried equilibrator headspace gas.  Gas flow is stopped prior to IR readings.

Drying Method:  The equilibrator headspace sample gas first goes through a glass

condenser cooled to ~ 5 oC.  The sample and standard gases pass through a short

column of magnesium perchlorate before reaching the analyzer.

Equilibrator (setup, size, flows):  The equilibrator is based on a design by R.

Weiss and was fabricated from a plexiglass housing with ~8 L water reservoir and

~16 L gaseous headspace.  Water flow rate is ~11 L/min.  Headspace recirculation

rate is ~200 ml/min.

Additional sensors:

Thermistor mounted in the bottom of the equilibrator.

Setra Barometer Model 370

YSI Model 600R thermosalinograph with temperature, salinity, and dissolved

oxygen sensors.  This TSG is mounted in the Hydro lab sink near the

equilibrator and the two are teed off the uncontaminated seawater feed.

The dissolved oxygen measurements are not reported in the final data file.

Narrative statement identifying measurement method for each required parameter:

CALCULATIONS:

The mixing ratios of ambient air and equilibrated headspace air are calculated

by fitting a second-order polynomial through the hourly averaged millivolt

response of the detector versus mixing ratios of the standards.  Mixing ratios

of dried equilibrated headspace and air are converted to fugacity of CO2 in

surface seawater and water saturated air in order to determine the fCO2.

For ambient air and equilibrator headspace, the fCO2a (or fCO2eq) is calculated

assuming 100% water vapor content:

fCO2eq = xCO2eq(P-pH2O)exp(B11+2*d12)P/RT

where fCO2eq is the fugacity in the equilibrator, pH2O is the water vapor

pressure at the sea surface temperature, P is the atmospheric pressure (in atm),

T is the SST or equilibrator temperature (in K) and R is the ideal gas constant

(82.057 cm^3·atm·deg^-1·mol^-1).  The exponential term is the fugacity correction

where B11 is the second virial coefficient of pure CO2

B11 = -1636.75 + 12.0408T - 0.032795T^2 + 3.16528E-5 T^3

and d12 = 57.7 - 0.118 T is the correction for an air-CO2 mixture in units of

cm^3·mol^-1 (Weiss, 1974).

The calculation for the fugacity at SST involves a temperature correction term

for the increase of fCO2 due to heating of the water from passing through the

pump and through 5 cm ID PVC tubing within the ship.  The empirical temperature

correction from equilibrator temperature to SST is:

fCO2(SST) = fCO2(eq) /

Exp ((Teq-SST) * [0.03107 – 2.7851E-4 * Teq – 1.8391E-3 * ln(fco2eq * 1.0E-6)])

where SST is sea surface temperature and Teq is the equilibrator temperature in

degrees °C.

Sampling Cycle:

The system runs on an hourly cycle during which 3 standard gases, 3 air

samples from the bow tower and 8 surface water samples (from the

equilibrator head space) are analyzed on the following schedule:

Mins. after hour		Sample

4 				Low Standard

8 				Mid Standard

12 				High Standard

16.5				Water

21 				Water

25.5				Water

30				Water

34				Air

38				Air

42				Air

46.5				Water

51				Water

55.5				Water

60				Water

NOTES ON DATA:

Columns have a default value of –999.99 in case of instrument malfunction,

erroneous readings or missing data.  Furthermore, if a suspicious xCO2 value,

pressure or temperature value is encountered, the fCO2 is not calculated.

Analytical Instrument Manufacturer/Model:

The Ron Brown system (version 2.6) was built by Craig Neill in 1999.  The analyzer 	is a LI-COR 6251 (analog output) infrared analyzer.

Standard Gases and Reference Gas:  The three standard gases came from CMDL

in Boulder and are directly traceable to the WMO scale.  While individual data

points above the high standard gas concentration or below the low standard gas

concentration may not be accurate, the general trends should be indicative of the

seawater chemistry.

Description of any additional environmental control:

The system is located in the Hydro Lab of the Ron Brown.  The room is

air-conditioned with little temperature fluctuation.

Resolution of measurement:

The resolution of the instrument is better than 0.1 ppm.

Estimated overall uncertainty of measurement:

The xCO2eq measurements are believed accurate to 0.1 ppm.  The

fCO2@SST measurements are believed to be precise to 0.2 ppm.

List of calibration gases used:

The standards used during the 2000 field season were:

STANDARD		TANK #	CONCENTRATION	VENDOR

Low			CA03334	298.47		CMDL

Low			CA03332	281.16		CMDL

Mid			CA03539	346.71		CMDL

Mid			CA03880	356.99		CMDL

High			CA01407	417.67		CMDL

High			CA03877	420.47		CMDL

High			N/A		524.99		CMDL

Traceability to an internationally recognized scale (including date/place of last calibration made):

All standards are obtained from NOAA/CMDL, now called the Global

Monitoring Division of the Earth Research Laboratory and are directly

traceable to WHO scale.

Uncertainty of assigned value of each calibration gas:

The uncertainty based on pre and post cruise calibrations is less

than 0.05 ppm.

Pressure/Temperature/Salinity:

For information about the ship’s thermosalinograph, contact Chief Survey Tech

Jonathan Shannahoff at jonathan.shannahoff@noaa.gov.

Units:

All xCO2 values are reported in parts per million (ppm) and fCO2 values

are reported in microatmospheres (uatm) assuming 100% humidity at the

equilibrator temperature.

Bibliography:

DOE (1994). Handbook of methods for the analysis of the various parameters of the

carbon dioxide system in sea water; version 2. DOE.

Feely, R. A., R. Wanninkhof, H. B. Milburn, C. E. Cosca, M. Stapp and P. P. Murphy

(1998). A new automated underway system for making high precision pCO2

measurements onboard research ships. Analytica Chim. Acta 377: 185-191.

Ho, D. T., R. Wanninkhof, J. Masters, R. A. Feely and C. E. Cosca (1997).

Measurement of underway fCO2 in the Eastern Equatorial Pacific on NOAA

ships BALDRIGE and DISCOVERER, NOAA data report ERL AOML-30, 52 pp.,

NTIS Springfield.

Wanninkhof, R. and K. Thoning (1993). Measurement of fugacity of CO2 in surface

water using continuous and discrete sampling methods. Mar. Chem. 44(2-4):

189-205.

Weiss, R. F. (1970). The solubility of nitrogen, oxygen and argon in water and

seawater. Deep-Sea Research 17: 721-735.

Weiss, R. F. (1974). Carbon dioxide in water and seawater: the solubility of

a non-ideal gas. Mar. Chem. 2: 203-215.

Weiss, R. F., R. A. Jahnke and C. D. Keeling (1982). Seasonal effects of

temperature and salinity on the partial pressure of CO2 in seawater.

Nature 300: 511-513.

Comments related to all 2000 data:

1.	xCO2 values outside the range of the standard gases (i.e. below the low standard or	above the high standard) are not as accurate as values within the range.  However,	the general trends should be indicative of the seawater chemistry.

2.	A new system built by Craig Neill (System 2.6) was installed for this field season.  	Side-by-side tests were conducted in the laboratory of systems 2.5 and 2.6.  The 	average difference in computed CO2 concentration between them was 0.2723 ppm with a 	standard deviation of 0.43575 (n = 289).  The main difference between the two 	systems is that the Licor infrared analyzer is now installed in a temperature 	controlled rack.  In addition, KNF pumps are now being used for both air and 	equilibrator headspace gas.

3.	The standard gases for the first 2 cruises were 298.47 ppm, 346.71 ppm, and 417.67 	ppm.  For the next 6 cruises the standards were 281.16 ppm, 356.99 ppm, and 420.47 	ppm.  For the last cruise the high standard was replaced with a 524.99 ppm standard 	due to the high CO2 concentrations in the study area.

Comments related to the individual legs:

RB200001:	1. Water values before March 18 at 1900 GMT were removed due to high 				variability in the readings.  This was caused by air leaks in the water 			lines.

2. The system was down from 2130 GMT on March 18 to 0300 GMT on March 				19 except for 1/2 hour while work was done on the plumbing connections.

3. The system was down from 2015 GMT to 2358 on March 20 for software 				modifications.

4. Data from 2045 GMT to 2400 on March 21 was discarded because system 				work corrupted the readings.

5. The system was down from 0430 to 0500 GMT on March 22 to replace a 				tube in the equilibrator.

RB200002:	1. Due to problems getting the system running, no data was collected 				before April 1.  Data in this file covers the period from April 1 at 				0000 GMT through April 4 at 0246 GMT.

2. The regulator on the mid standard tank failed and there was no gas 				flow in that phase for the entire cruise.  Examination of the Licor 				voltage response during the previous cruise showed a relationship 				between mid standard voltage response and voltage response in the low 				and high standard phases.  This relationship is (VM - VL)/(VH - VL) = 				0.421 where VL, VM, and VH are voltages in the low, mid and high 					standard phases respectively.  Using this equation, a synthetic mid 				standard voltage was created and applied to the previous cruise.  The 				difference between using the real and synthetic voltages for the 					previous cruise was less than 0.5 ppm.  All concentrations for the 				current cruise were computed using this synthetic mid standard voltage.

RB200003:	1. A high cyclical variation in air xCO2 values caused me to suspect a 				leak.  I therefore replaced all air xCO2 values with 376.00, which I 				derived from climatology records provided by CMDL in Boulder.  I used 				values from Cold Harbor, Alaska and Cape Meares, Oregon from 1993 to 				1997.  For each year I averaged all values in the period from late May 				to early June and plotted them by year.  I used a linear fit to 					extrapolate each station's values to the year 2000 to arrive at the 				value of 376.00.  This matches fairly closely with minimum air xCO2 				values observed during the cruise which fell in the range of 374.50 to 				377.00.

RB200004:	1. Due to high variation in air xCO2 values I replaced all air xCO2 				values with 375.00, which I derived from climatology records provided 				by CMDL in Boulder.  I used values from the station in Cold Harbor, 				Alaska from 1994 to 1998.  For each year I averaged all readings taken 				between June 17 and June 27 and plotted them by year.  I used a linear 				fit to extrapolate to the year 2000 to arrive at the value of 375.00.

RB200005:	1. Due to high variation in air xCO2 values I replaced all air xCO2 				values with 366.50, which I derived from climatology records provided 				by CMDL in Boulder.  I used values from the station in Cape Meares, OR 				from 1989 to 1997.  I removed values for 1992 and 1993 because they 				ruined the linear curve fit.  There were no good values for 1995 and 				1996.  For each year I averaged all readings taken between June 29 and 				July 18 and plotted them by year.  I used a linear fit to extrapolate 				to the year 2000 to arrive at the value of 366.50.  This value is in 				close agreement with the minimum air values observed during the cruise.

RB200006:	1. Due to high variation in air xCO2 values I replaced all air xCO2 				values with 363.00, which I derived from climatology records provided 				by CMDL in Boulder.  I used values from the station in Cape Meares, OR 				from 1989 to 1997.  There were no good values for 1993 and 1995.  For 				each year I averaged all readings taken between July 20 and August 4 				and plotted them by year.  I used a linear fit to extrapolate to the 				year 2000 to arrive at the value of 363.00.  This value is in close 				agreement with the minimum air values observed during the cruise.

2. On July 29 at approximately 1800 GMT the difference between the 				equilibrator temperature (EqT) and SST drops about 0.11 degrees C (from 			~ 0.42 to ~ 0.31) and stays that way for the remainder of the cruise.  				EqT is measured at the equilibrator in the Hydro Lab and SST is 					measured at the seawater intake in the bow of the ship.  For SST in the 			range of 16-17 degrees C, the normal difference is about 0.4 degrees.

RB200007:	No problems of note.

RB200008:	No problems of note.

RB200009:	No problems of note.