R/V Ronald H. Brown METADATA - 2005

Class of Data:  Surface ocean and atmospheric carbon dioxide concentrations

Dataset Identifier:  R/V Ronald H. Brown

One File:  RHB2005

Statement of how to cite dataset:

Ron Brown website: http://www.aoml.noaa.gov/ocd/gcc/rvbrown_data2005.php

These data are made freely available to the public and the scientific community

in the belief that their wide dissemination will lead to greater understanding

and new scientific insights.  The availability of these data does not

constitute publication of the data.  We rely on the ethics and integrity of

the user to assure that AOML receives fair credit for our work.  Please send

manuscripts using this data to AOML for review before they are submitted

for publication so we can insure that the quality and limitations of the data

are accurately represented.

Measurement platform identifier:  NOAA research vessel Ronald H. Brown (R104)

Cruise Information:

The Ron Brown conducted 6 major cruises in the Atlantic and eastern Pacific

Oceans for a total of 11 legs.

Project Information:

The system was operated by personnel from AOML or PMEL (Pacific Marine 	Environmental Laboratory) or by the Ron Brown’s Chief Survey Tech, Jonathan 	Shannahoff.  The work was sponsored by the Underway pCO2 on Ships project of the 	NOAA climate program.

Scientist responsible for technical quality of dataset:

Rik Wanninkhof

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Rik.Wanninkhof@noaa.gov

Contact person for this dataset:

Bob Castle

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Robert.Castle@noaa.gov

Timestamp for initial submission of dataset:  11/18/09

Timestamp for the most recent update of dataset:  11/18/09

Timestamp period the dataset refers to: 12/29/2004 – 11/25/2005

Geographic area the dataset refers to:

65 S to 45 N

125 W to 20 W

2005 Cruises:

RB200501T –	CLIVAR A16S Transit

Valparaiso, Chile to Punta Arenas, Chile

December 29, 2004 to January 5, 2005

Chief Scientist – Kevin Sullivan

Operator – Jonathan Shannahoff

RB200501A -	CLIVAR A16S

Punta Arenas, Chile to Fortaleza, Brazil

January 11, 2005 to February 22, 2005

Chief Scientist – Rik Wanninkhof

Operator – Jonathan Shannahoff

RB200502T -	NTAS Moorings Transit

Fortaleza, Brazil to Bridgetown, Barbados

February 28, 2005 to March 6, 2005

Chief Scientist – N/A

Operator – Jonathan Shannahoff

RB200502A -	NTAS Moorings

Bridgetown, Barbados to Charleston, SC

March 10, 2005 to March 18, 2005

Chief Scientist – Al Plueddemann

Operator – Jonathan Shannahoff

RB200503SB-	Sea Beam Patch Testing

Charleston, SC to Woods Hole, MA

July 12, 2005 to July 15, 2005

Chief Scientist – N/A

Operator – Jonathan Shannahoff

RB200503A -	Ocean Explorations Lost City

Woods Hole, MA, to Punta Delgada, Azores

July 18, 2005 to August 4, 2005

Chief Scientist – Robert Ballard

Operator – Jonathan Shannahoff

RB200503B -	Ocean Explorations Stepping Stone

Punta Delgada, Azores to Woods Hole, MA

August 10, 2005 to September 3, 2005

Chief Scientist – Les Watling

Operator – Jonathan Shannahoff

RB200504T -	Western Boundary Time Series Transit

Woods Hole to Charleston, SC

September 6, 2005 to September 9, 2005

Chief Scientist – N/A

Operator – Jonathan Shannahoff

RB200504 -	Western Boundary Time Series

Charleston, SC to Miami, FL

September 11, 2005 to September 24, 2005

Chief Scientist – Chris Meinen

Operator – Jonathan Shannahoff

RB200505 -	Stratus Mooring

Rodman, Panama to Arica, Chile

October 4, 2005 to October 20, 2005

Chief Scientist – Robert Weller

Operator – Jonathan Shannahoff

RB200506 -	TAO

Arica, Chile to Rodman, Panama

October 27, 2005 to November 25, 2005

Chief Scientist – Andy Shepherd

Operator – Jonathan Shannahoff

List of variables included in this dataset:

COLUMN	HEADER			EXPLANATION

1.		GROUP/SHIP:			AOML_Brown for all underway data from the Ron Brown.

2.		CRUISE_DESIGNATION:	Cruise ID (e.g., RBYYYYnn where RB = Ron Brown, YYYY

= the four digit year, and nn = the cruise number for

that year).

3.		JD_GMT:			Decimal year day.

4.		DATE_DDMMYYYY:		GMT date.  The date format has been changed to comply

with the IOCCP recommendations.

5.		TIME_HH:MM:SS:		GMT time.

6.		LAT_DEC_DEGREE:		Latitude in decimal degrees (negative values are in 							the southern hemisphere).

7.		LONG_DEC_DEGREE:		Longitude in decimal degrees (negative values are in 							the western hemisphere).

8.		xCO2W_PPM:			Mole fraction of CO2 (dry) in the equilibrator

headspace at equilibrator temperature (Teq) in parts

per million.

9.		xCO2A_PPM:			Mole fraction of CO2 in air in parts per million.

10.		PRES_EQUIL_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).

11.		PRES_SEALEVEL_hPa:	Barometric pressure corrected to sea level from the 							ship’s barometer in hectopascals (1 hectopascal = 1 							millibar).

12.		EqTEMP_C:			Temperature in equilibrator water in degrees

centigade.  Temperature in equilibrator measured with

a calibrated thermistor.

13.		SST(TSG)_C:			Temperature from the ship's thermosalinograph in

degrees centigrade.

14.		SAL(TSG)_PERMIL:		Salinity from the ship's thermosalinograph on the

Practical Salinity Scale.

15.		WATER_FLOW_L/MIN:		Water flow rate through the equilibrator in liters 							per minute.

16.		GAS_FLOW_IR_ML/MIN:	Gas flow through the sample cell of the Licor IR 							analyzer in milliliters per minute.

17.		TEMP_IR_C:			Temperature in the Licor sample cell in degrees 							centigrade.

18.		PRES_IR_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).  The Licor in this system 							does not include a pressure sensor so this field is 							the same as # 10 above.

19.		SHIP_HEADING_TRUE_DEGREE:	Ship’s heading in true degrees from the ship’s 								scientific computing system.

20.		SHIP_SPEED_KNOT:		Ship’s speed in knots from the ship’s scientific 							computing system.

21.		WIND_DIR_REL_DEGREE:	Relative wind direction in degrees from the ship’s 							scientific computing system.

22.		WIND_SPEED_REL_M/S:	Relative wind speed in meters per second from the 							ship’s scientific computing system.

23.		fCO2W@SST_uatm:		Fugacity of CO2 in sea water in microatmospheres

calculated as outlined below.

24.		QC_FLAG_WATER:		Quality control flag for fCO2W@SST measurement.  2 = 							good, 3 = questionable, 4 = bad.

25.		fCO2A_uATM:			Fugacity of CO2 in air in microatmospheres

calculated as outlined below.

24.		QC_FLAG_AIR:		Quality control flag for fCO2A measurement.  2 = 							good, 3 = questionable, 4 = bad.

27.		dfCO2_uATM:			Sea water fCO2 - air fCO2 in microatmospheres.  This

uses the average air value for the current hour.

28.		FLUORO_uG/l:		Measurement from the ship’s Turner 10AU fluorometer 							in micrograms per liter.

29.		WIND_SPEED_TRUE_M/S:	True wind speed in meters per second from the 								ship’s scientific computing system.

30.		WIND_DIR_TRUE_DEGREE:	True wind direction in degrees from the ship’s 								scientific computing system.

31.		AIR_TEMP_C:			Outside air temperature from the ship’s scientific 							computing system.

The following fields have been QC'ed by the CO2 group:

GROUP/SHIP

CRUISE_DESIGNATION

JD_GMT

DATE_DDMMYYYY

TIME_HH:MM:SS

LAT_DEC_DEGREE

LONG_DEC_DEGREE

xCO2W_PPM

xCO2A_PPM

EqTEMP_C

PRES_EQUIL_hPa

WATER_FLOW_L/MIN

GAS_FLOW_L/MIN

TEMP_IR_C

PRES_IR_hPa

fCO2W@SST_uatm

fCO2A_uATM

dfCO2_uatm

The following fields are from the ship's onboard systems and the quality of this

data cannot be verified:

SST(TSG)_C

Sal(TSG)_Permil

PRES_SEALEVEL_hPa

SHIP_HEADING_TRUE_DEGREE

SHIP_SPEED_KNOT

WIND_DIR_REL_DEGREE

WIND_SPEED_REL_M/S

FLUORO_uG/l

WIND_SPEED_TRUE_M/S

WIND_DIR_TRUE_DEGREE

AIR_TEMP_C

Narrative description of system design:

CO2 ANALYTICAL SYSTEM:

The concentration of carbon dioxide (CO2) in surface ocean water is determined

by measuring the concentration of CO2 in gas that is in contact with the water.

Surface water is pumped ~ 100 m through 7/8" Teflon tubing from an inlet

in the ship's bow to the equilibration chamber.  Water comes from the bow

intake ~4.2 m below the water line and the TSG is located close to the inlet.

When the SST is below about 20 oC, friction in the pipes and from the pump cause

heating and the Teq is higher than SST.  When the SST is higher than about 25 oC,

the ship’s air conditioning cools the water and the Teq is lower than SST.

The equilibration chamber has an enclosed volume of gas, or headspace, and a pool

of seawater that continuously overflows to a drain.  As the water flows through the

chamber, the dissolved gases (like CO2) partition between the water and the

headspace.  At equilibrium, the ratio of CO2 in the water and in the headspace is

influenced most by temperature, and that relationship is known.  By measuring

the concentration of CO2 in the headspace and the temperature in the chamber,

the partial pressure (or fugacity) of CO2 in the surface water can be calculated.

INSTRUMENT DESCRIPTION

The general principle of instrumental design can be found in Wanninkhof and Thoning

(1993), Ho et al. (1995), and Feely et al. (1999).  The concentration of CO2 in the

headspace gas is measured using the adsorption of infrared (IR) radiation, which

results from changes in the rotational and vibrational energy state of the CO2

molecule.  The LI-COR detector passes IR radiation through two 6" cells.  The

reference cell is flushed with a gas of known CO2 concentration.  The sample cell

is flushed with the headspace gas.  A vacuum-sealed, heated filament is the

broadband IR source.  The IR radiation alternates between the two cells via a

chopping shutter disc. An optical filter selects an adsorption band specific

for CO2 (4.26 micron) to reach the detector.  The solid state (lead selenide)

detector is kept at -12 degrees °C for excellent stability and low signal

noise (less than 0.2 ppm).

Several steps are taken to reduce interferences and to increase the accuracy

of the measurements.  After the equilibration chamber, the headspace travels

through a drying trap to remove water vapor.  During each analysis, the

headspace gas is compared to a reference gas of known concentration.  To

improve the accuracy of the measurements, three different gaseous standards

for CO2 are analyzed once an hour instead of the headspace gas.

Analyzer:  LI-COR 6251 (analog output) infrared (IR) analyzer.

Method of Analysis:  Differential analyses relative to the low standard.  Measures

dried equilibrator headspace gas.  Gas flow is stopped prior to IR readings.

Drying Method:  The equilibrator headspace sample gas first goes through a glass

condenser cooled to ~ 5 oC.  The sample and standard gases pass through a short

column of magnesium perchlorate before reaching the analyzer.

Equilibrator (setup, size, flows):  The equilibrator is based on a design by R.

Weiss and was fabricated from a plexiglass housing with ~8 L water reservoir and

~16 L gaseous headspace.  Water flow rate is ~11 L/min.  Headspace recirculation

rate is ~200 ml/min.

Additional sensors:

Thermistor mounted in the bottom of the equilibrator.

Setra Barometer Model 370

YSI Model 600R thermosalinograph with temperature, salinity, and dissolved

oxygen sensors.  This TSG is mounted in the Hydro lab sink near the

equilibrator and the two are teed off the uncontaminated seawater feed.

The dissolved oxygen measurements are not reported in the final data file.

Narrative statement identifying measurement method for each required parameter:

CALCULATIONS:

The mixing ratios of ambient air and equilibrated headspace air are calculated

by fitting a second-order polynomial through the hourly averaged millivolt

response of the detector versus mixing ratios of the standards.  Mixing ratios

of dried equilibrated headspace and air are converted to fugacity of CO2 in

surface seawater and water saturated air in order to determine the fCO2.

For ambient air and equilibrator headspace, the fCO2a (or fCO2eq) is calculated

assuming 100% water vapor content:

fCO2eq = xCO2eq(P-pH2O)exp(B11+2*d12)P/RT

where fCO2eq is the fugacity in the equilibrator, pH2O is the water vapor

pressure at the sea surface temperature, P is the atmospheric pressure (in atm),

T is the SST or equilibrator temperature (in K) and R is the ideal gas constant

(82.057 cm^3·atm·deg^-1·mol^-1).  The exponential term is the fugacity correction

where B11 is the second virial coefficient of pure CO2

B11 = -1636.75 + 12.0408T - 0.032795T^2 + 3.16528E-5 T^3

and d12 = 57.7 - 0.118 T is the correction for an air-CO2 mixture in units of

cm^3·mol^-1 (Weiss, 1974).

The calculation for the fugacity at SST involves a temperature correction term

for the increase of fCO2 due to heating of the water from passing through the

pump and through 5 cm ID PVC tubing within the ship.  The empirical temperature

correction from equilibrator temperature to SST is:

fCO2(SST) = fCO2(eq) /

Exp ((Teq-SST) * [0.03107 – 2.7851E-4 * Teq – 1.8391E-3 * ln(fco2eq * 1.0E-6)])

where SST is sea surface temperature and Teq is the equilibrator temperature in

degrees °C.

Sampling Cycle:

The system runs on an hourly cycle during which 3 standard gases, 3 air

samples from the bow tower and 8 surface water samples (from the

equilibrator head space) are analyzed on the following schedule:

Mins. after hour		Sample

4 				Low Standard

8 				Mid Standard

12 				High Standard

16.5				Water

21 				Water

25.5				Water

30				Water

34				Air

38				Air

42				Air

46.5				Water

51				Water

55.5				Water

60				Water

NOTES ON DATA:

Columns have a default value of –999.99 in case of instrument malfunction,

erroneous readings or missing data.  Furthermore, if a suspicious xCO2 value,

pressure or temperature value is encountered, the fCO2 is not calculated.

Analytical Instrument Manufacturer/Model:

The Ron Brown system (version 2.6) was built by Craig Neill in 1999.  The analyzer 	is a LI-COR 6251 (analog output) infrared analyzer.

Standard Gases and Reference Gas:  The three standard gases came from CMDL

in Boulder and are directly traceable to the WMO scale.  While individual data

points above the high standard gas concentration or below the low standard gas

concentration may not be accurate, the general trends should be indicative of the

seawater chemistry.

Description of any additional environmental control:

The system is located in the Hydro Lab of the Ron Brown.  The room is

air-conditioned with little temperature fluctuation.

Resolution of measurement:

The resolution of the instrument is better than 0.1 ppm.

Estimated overall uncertainty of measurement:

The xCO2eq measurements are believed accurate to 0.1 ppm.  The

fCO2@SST measurements are believed to be precise to 0.2 ppm.

List of calibration gases used:

The standards used during the 2005 field season were:

STANDARD		TANK #	CONCENTRATION	VENDOR

Low			CA05395	315.25		ESRL

Low			CA05760	303.23		ESRL

Mid			CA05398	370.90		ESRL

Mid,high		CA05344	411.42		ESRL

High			CC71588	531.98		ESRL

Traceability to an internationally recognized scale (including date/place of last calibration made):

All standards are obtained from NOAA/CMDL, now called the Global

Monitoring Division of the Earth System Research Laboratory and are directly

traceable to WMO scale.

Uncertainty of assigned value of each calibration gas:

The uncertainty based on pre and post cruise calibrations is less

than 0.05 ppm.

Pressure/Temperature/Salinity:

For information about the ship’s thermosalinograph or other sensors, contact Chief 	Survey Tech Jonathan Shannahoff at jonathan.shannahoff@noaa.gov.

Units:

All xCO2 values are reported in parts per million (ppm) and fCO2 values

are reported in microatmospheres (uatm) assuming 100% humidity at the

equilibrator temperature.

Bibliography:

DOE (1994). Handbook of methods for the analysis of the various parameters of the

carbon dioxide system in sea water; version 2. DOE.

Feely, R. A., R. Wanninkhof, H. B. Milburn, C. E. Cosca, M. Stapp and P. P. Murphy

(1998). A new automated underway system for making high precision pCO2

measurements onboard research ships. Analytica Chim. Acta 377: 185-191.

Ho, D. T., R. Wanninkhof, J. Masters, R. A. Feely and C. E. Cosca (1997).

Measurement of underway fCO2 in the Eastern Equatorial Pacific on NOAA

ships BALDRIGE and DISCOVERER, NOAA data report ERL AOML-30, 52 pp.,

NTIS Springfield.

Wanninkhof, R. and K. Thoning (1993). Measurement of fugacity of CO2 in surface

water using continuous and discrete sampling methods. Mar. Chem. 44(2-4):

189-205.

Weiss, R. F. (1970). The solubility of nitrogen, oxygen and argon in water and

seawater. Deep-Sea Research 17: 721-735.

Weiss, R. F. (1974). Carbon dioxide in water and seawater: the solubility of

a non-ideal gas. Mar. Chem. 2: 203-215.

Weiss, R. F., R. A. Jahnke and C. D. Keeling (1982). Seasonal effects of

temperature and salinity on the partial pressure of CO2 in seawater.

Nature 300: 511-513.

Comments related to all 2005 data:

1.	xCO2 values outside the range of the standard gases (i.e. below the low standard or	above the high standard) are not as accurate as values within the range.  However,	the general trends should be indicative of the seawater chemistry.

2.	The standard gases for the first 4 cruises (RB200501T-RB200502A) were 315.25 ppm, 	411.42 ppm, and 531.98 ppm.  The standards for the last 7 cruises (RB200503SB-	RB200506) were 303.23 ppm, 370.90 ppm, and 411.42 ppm.

Comments related to the individual legs:

RB200501T:	1. Extremely high pCO2 values were encountered in the first two days of 			the cruise.  These are apparently due to upwelling and have been left 				in the data file since there is no indication of problems with the 				system.

2. Air values have been removed in three periods:  from 0740 to 1550 on 			December 30, from 0840 to 1550 on January 2, and from 2345 on January 3 			to 0940 on January 4.  In the last two cases these high values are 				likely due to stack gas contamination.  In the first case, it is 					possible that they are due to shore-based sources.  Some of these have 				been left in and given QC flags of 3.

3. The Seabird TSG was not working during this leg and the temperature 				and salinity from it should not be considered valid.  Instead, the 				temperature and salinity from the ship's TSG or the equilibrator 					temperature should be used.

RB200501A:	1. AC power problems occurred on Jan. 14 from 1344 to 1830, on Jan. 15 				from 2000 to 2300, and on Jan. 16 from 144 to 1544.  During these 				periods, most water and air values have been removed.

2. Also during the periods of power problems, the equilibrator 					thermistor gave bad readings.  A relationship was derived between the 				Seabird temperature and the equilibrator temperature (EqT = (0.0129 * 				SBT) - 0.2932, R^2 = .9282) and was used to replace the bad 					equilibrator temperatures.

RB200502T:	No problems of note.

RB200502A:	No problems of note.

RB200503SB:	1. There was no water flow from the beginning of the cruise until July 				12 at 1230.  All water values were removed through July 12 at 1330.

RB200503A:	1. A severe drop in gas flow in the water phase occurred on July 25 at 				around 0600.  This was probably caused by an obstruction in the gas 				line.  About half of  the water values from July 25 at 0600 until July 				27 at 1200 were removed and all water values after noon on July 27 were 			removed.

RB200503B:	1. There was insufficient gas flow in the water phase from the 					beginning of the cruise until about 2245 on August 18.  Almost all 				water values in this time period have been removed.

RB200504T:	No problems of note.

RB200504A:	No problems of note.

RB200505:	1. There was no gas flow in the water phase from October 5 at 1357 to 				October 6 at 1253.  All water values during this time have been 					removed.

2. The salinity readings from the ship's TSG were bad from October 13 				at 0019 to October 15 at 1703.  They were replaced with values from the 			Seabird TSG near the underway system using the equation SCS sal = SBE 				sal - 0.16, which was derived from the period when both sensors were 				working correctly.

RB200506:	1. There was a problem with the ship's salinity sensor and all values 				before October 28 at 1645 and after November 21 at 1530 were replaced.  			The new values were equal to the salinity from the Seabird Micro TSG - 				0.077 where 0.077 was the average difference between the ship's TSG and 			the Seabird Micro TSG when the ship's TSG was

giving good values.

2. The gas flow in the water phase dropped to near zero on November 10 				from 0420 to 1304 and from November 22 at 0501 to November 23 at 1327.  			Almost all water values during these time periods have been removed.