R/V Ronald H. Brown METADATA - 2007

Class of Data:  Surface ocean and atmospheric carbon dioxide concentrations

Dataset Identifier:  R/V Ronald H. Brown

One File:  RHB2007

Statement of how to cite dataset:

Ron Brown website: http://www.aoml.noaa.gov/ocd/gcc/rvbrown_data2007.php

These data are made freely available to the public and the scientific community

in the belief that their wide dissemination will lead to greater understanding

and new scientific insights.  The availability of these data does not

constitute publication of the data.  We rely on the ethics and integrity of

the user to assure that AOML receives fair credit for our work.  Please send

manuscripts using this data to AOML for review before they are submitted

for publication so we can insure that the quality and limitations of the data

are accurately represented.

Measurement platform identifier:  NOAA research vessel Ronald H. Brown (R104)

Cruise Information:

The Ron Brown conducted 11 major cruises in the Atlantic and eastern Pacific

Oceans for a total of 11 legs.

Project Information:

The system was operated by personnel from AOML or PMEL (Pacific Marine 	Environmental Laboratory) or by the Ron Brown’s Chief Survey Tech, Jonathan 	Shannahoff.  The work was sponsored by the Underway pCO2 on Ships project of the 	NOAA climate program.

Scientist responsible for technical quality of dataset:

Rik Wanninkhof

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Rik.Wanninkhof@noaa.gov

Contact person for this dataset:

Bob Castle

NOAA/AOML/Ocean Chemistry Division

4301 Rickenbacker Causeway

Miami, Florida 33149

Robert.Castle@noaa.gov

Timestamp for initial submission of dataset:  12/03/09

Timestamp for the most recent update of dataset:  12/03/09

Timestamp period the dataset refers to: 03/23/2007 – 01/18/2008

Geographic area the dataset refers to:

30 S to 45 N

120 W to 20 W

2007 Cruises:

RB200701 –	Western Boundary Time Series

Charleston, SC to San Juan, Puerto Rico

March 23, 2007 to April 10, 2007

Chief Scientist – Molly Baringer

Operator – Jonathan Shannahoff

RB200702 -	NTAS

San Juan, Puerto Rico to Bridgetown, Barbados

April 16, 2007 to April 28, 2007

Chief Scientist – Al Plueddemann

Operator – Jonathan Shannahoff

RB200703 -	PNE

Bridgetown, Barbados to Ft. Lauderdale, FL

May 5, 2007 to May 29, 2007

Chief Scientist – Rick Lumpkin

Operator – Jonathan Shannahoff

RB200704 -	Ocean Explorations

Ft. Lauderdale, FL to Galveston, TX

June 4, 2007 to July 6, 2007

Chief Scientist – Charles Fisher

Operator – Jonathan Shannahoff

RB200705-	GOMECC

Galveston, TX to Boston, MA

July 10, 2007 to August 4, 2007

Chief Scientist – Tsung-Hung Peng

Operator – Jonathan Shannahoff

RB200706 -	NURP

Boston, MA, to Buzzard’s Bay, MA

August 8, 2007 to August 23, 2007

Chief Scientist – Peter Rona

Operator – Jonathan Shannahoff

RB200707 -	DART

Composition and Climate Study Leg A

Buzzard’s Bay, MA to Charleston, SC

August 26, 2007 to September 5, 2007

Chief Scientist – Mike Brewer

Operator – Jonathan Shannahoff

RB200708 -	Texas Air Quality Study/Gulf of Mexico Atmospheric

Composition and Climate Study Leg A

Charleston, SC to Charleston, SC

September 11, 2007 to September 22, 2007

Chief Scientist – Chris Meinen

Operator – Jonathan Shannahoff

RB200709 -	Stratus

Rodman, Panama to Galapagos Islands

October 16, 2007 to November 7, 2007

Chief Scientist – Bob Weller

Operator – Dana Mancinelli

RB200710 -	TAO

Galapagos Islands to San Diego, CA

November 11, 2007 to November 22, 2007

Chief Scientist – Robert Weller

Operator – Jonathan Shannahoff

RB200711 -	CLIVAR P18 Leg 1

San Diego, CA to Rodman, Panama

December 15, 2007 to January 18, 2007

Chief Scientist – John Bullister

Operator – Jonathan Shannahoff

List of variables included in this dataset:

COLUMN	HEADER			EXPLANATION

1.		GROUP/SHIP:			AOML_Brown for all underway data from the Ron Brown.

2.		CRUISE_DESIGNATION:	Cruise ID (e.g., RBYYYYnn where RB = Ron Brown, YYYY

= the four digit year, and nn = the cruise number for

that year).

3.		JD_GMT:			Decimal year day.

4.		DATE_DDMMYYYY:		GMT date.  The date format has been changed to comply

with the IOCCP recommendations.

5.		TIME_HH:MM:SS:		GMT time.

6.		LAT_DEC_DEGREE:		Latitude in decimal degrees (negative values are in 							the southern hemisphere).

7.		LONG_DEC_DEGREE:		Longitude in decimal degrees (negative values are in 							the western hemisphere).

8.		xCO2W_PPM:			Mole fraction of CO2 (dry) in the equilibrator

headspace at equilibrator temperature (Teq) in parts

per million.

9.		xCO2A_PPM:			Mole fraction of CO2 in air in parts per million.

10.		PRES_EQUIL_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).

11.		PRES_SEALEVEL_hPa:	Barometric pressure corrected to sea level from the 							ship’s barometer in hectopascals (1 hectopascal = 1 							millibar).

12.		EqTEMP_C:			Temperature in equilibrator water in degrees

centigade.  Temperature in equilibrator measured with

a calibrated thermistor.

13.		SST(TSG)_C:			Temperature from the ship's thermosalinograph in

degrees centigrade.

14.		SAL(TSG)_PERMIL:		Salinity from the ship's thermosalinograph on the

Practical Salinity Scale.

15.		WATER_FLOW_L/MIN:		Water flow rate through the equilibrator in liters 							per minute.

16.		GAS_FLOW_IR_ML/MIN:	Gas flow through the sample cell of the Licor IR 							analyzer in milliliters per minute.

17.		TEMP_IR_C:			Temperature in the Licor sample cell in degrees 							centigrade.

18.		PRES_IR_hPa:		Barometric pressure in the lab in hectopascals (1

hectopascal = 1 millibar).  The Licor in this system 							does not include a pressure sensor so this field is 							the same as # 10 above.

19.		SHIP_HEADING_TRUE_DEGREE:	Ship’s heading in true degrees from the ship’s 								scientific computing system.

20.		SHIP_SPEED_KNOT:		Ship’s speed in knots from the ship’s scientific 							computing system.

21.		WIND_DIR_REL_DEGREE:	Relative wind direction in degrees from the ship’s 							scientific computing system.

22.		WIND_SPEED_REL_M/S:	Relative wind speed in meters per second from the 							ship’s scientific computing system.

23.		fCO2W@SST_uatm:		Fugacity of CO2 in sea water in microatmospheres

calculated as outlined below.

24.		QC_FLAG_WATER:		Quality control flag for fCO2W@SST measurement.  2 = 							good, 3 = questionable, 4 = bad.

25.		fCO2A_uATM:			Fugacity of CO2 in air in microatmospheres

calculated as outlined below.

24.		QC_FLAG_AIR:		Quality control flag for fCO2A measurement.  2 = 							good, 3 = questionable, 4 = bad.

27.		dfCO2_uATM:			Sea water fCO2 - air fCO2 in microatmospheres.  This

uses the average air value for the current hour.

28.		FLUORO_uG/l:		Measurement from the ship’s Turner 10AU fluorometer 							in micrograms per liter.

29.		WIND_SPEED_TRUE_M/S:	True wind speed in meters per second from the 								ship’s scientific computing system.

30.		WIND_DIR_TRUE_DEGREE:	True wind direction in degrees from the ship’s 								scientific computing system.

31.		AIR_TEMP_C:			Outside air temperature from the ship’s scientific 							computing system.

The following fields have been QC'ed by the CO2 group:

GROUP/SHIP

CRUISE_DESIGNATION

JD_GMT

DATE_DDMMYYYY

TIME_HH:MM:SS

LAT_DEC_DEGREE

LONG_DEC_DEGREE

xCO2W_PPM

xCO2A_PPM

EqTEMP_C

PRES_EQUIL_hPa

WATER_FLOW_L/MIN

GAS_FLOW_L/MIN

TEMP_IR_C

PRES_IR_hPa

fCO2W@SST_uatm

fCO2A_uATM

dfCO2_uatm

The following fields are from the ship's onboard systems and the quality of this

data cannot be verified:

SST(TSG)_C

Sal(TSG)_Permil

PRES_SEALEVEL_hPa

SHIP_HEADING_TRUE_DEGREE

SHIP_SPEED_KNOT

WIND_DIR_REL_DEGREE

WIND_SPEED_REL_M/S

FLUORO_uG/l

WIND_SPEED_TRUE_M/S

WIND_DIR_TRUE_DEGREE

AIR_TEMP_C

Narrative description of system design:

CO2 ANALYTICAL SYSTEM:

The concentration of carbon dioxide (CO2) in surface ocean water is determined

by measuring the concentration of CO2 in gas that is in contact with the water.

Surface water is pumped ~ 100 m through 7/8" Teflon tubing from an inlet

in the ship's bow to the equilibration chamber.  Water comes from the bow

intake ~4.2 m below the water line and the TSG is located close to the inlet.

When the SST is below about 20 oC, friction in the pipes and from the pump cause

heating and the Teq is higher than SST.  When the SST is higher than about 25 oC,

the ship’s air conditioning cools the water and the Teq is lower than SST.

The equilibration chamber has an enclosed volume of gas, or headspace, and a pool

of seawater that continuously overflows to a drain.  As the water flows through the

chamber, the dissolved gases (like CO2) partition between the water and the

headspace.  At equilibrium, the ratio of CO2 in the water and in the headspace is

influenced most by temperature, and that relationship is known.  By measuring

the concentration of CO2 in the headspace and the temperature in the chamber,

the partial pressure (or fugacity) of CO2 in the surface water can be calculated.

INSTRUMENT DESCRIPTION

The general principle of instrumental design can be found in Wanninkhof and Thoning

(1993), Ho et al. (1995), and Feely et al. (1999).  The concentration of CO2 in the

headspace gas is measured using the adsorption of infrared (IR) radiation, which

results from changes in the rotational and vibrational energy state of the CO2

molecule.  The LI-COR detector passes IR radiation through two 6" cells.  The

reference cell is flushed with a gas of known CO2 concentration.  The sample cell

is flushed with the headspace gas.  A vacuum-sealed, heated filament is the

broadband IR source.  The IR radiation alternates between the two cells via a

chopping shutter disc. An optical filter selects an adsorption band specific

for CO2 (4.26 micron) to reach the detector.  The solid state (lead selenide)

detector is kept at -12 degrees °C for excellent stability and low signal

noise (less than 0.2 ppm).

Several steps are taken to reduce interferences and to increase the accuracy

of the measurements.  After the equilibration chamber, the headspace travels

through a drying trap to remove water vapor.  During each analysis, the

headspace gas is compared to a reference gas of known concentration.  To

improve the accuracy of the measurements, three different gaseous standards

for CO2 are analyzed once an hour instead of the headspace gas.

Analyzer:  LI-COR 6251 (analog output) infrared (IR) analyzer.

Method of Analysis:  Differential analyses relative to the low standard.  Measures

dried equilibrator headspace gas.  Gas flow is stopped prior to IR readings.

Drying Method:  The equilibrator headspace sample gas first goes through a glass

condenser cooled to ~ 5 oC.  The sample and standard gases pass through a short

column of magnesium perchlorate before reaching the analyzer.

Equilibrator (setup, size, flows):  The equilibrator is based on a design by R.

Weiss and was fabricated from a plexiglass housing with ~8 L water reservoir and

~16 L gaseous headspace.  Water flow rate is ~11 L/min.  Headspace recirculation

rate is ~200 ml/min.

Additional sensors:

Thermistor mounted in the bottom of the equilibrator.

Setra Barometer Model 370

Seabird Micro TSG Model SBE-45 with temperature and salinity sensors.  This TSG is 			mounted in the Hydro lab sink near the equilibrator and the two are teed off 			the uncontaminated seawater feed.

Narrative statement identifying measurement method for each required parameter:

CALCULATIONS:

The mixing ratios of ambient air and equilibrated headspace air are calculated

by fitting a second-order polynomial through the hourly averaged millivolt

response of the detector versus mixing ratios of the standards.  Mixing ratios

of dried equilibrated headspace and air are converted to fugacity of CO2 in

surface seawater and water saturated air in order to determine the fCO2.

For ambient air and equilibrator headspace, the fCO2a (or fCO2eq) is calculated

assuming 100% water vapor content:

fCO2eq = xCO2eq(P-pH2O)exp(B11+2*d12)P/RT

where fCO2eq is the fugacity in the equilibrator, pH2O is the water vapor

pressure at the sea surface temperature, P is the atmospheric pressure (in atm),

T is the SST or equilibrator temperature (in K) and R is the ideal gas constant

(82.057 cm^3·atm·deg^-1·mol^-1).  The exponential term is the fugacity correction

where B11 is the second virial coefficient of pure CO2

B11 = -1636.75 + 12.0408T - 0.032795T^2 + 3.16528E-5 T^3

and d12 = 57.7 - 0.118 T is the correction for an air-CO2 mixture in units of

cm^3·mol^-1 (Weiss, 1974).

The calculation for the fugacity at SST involves a temperature correction term

for the increase of fCO2 due to heating of the water from passing through the

pump and through 5 cm ID PVC tubing within the ship.  The empirical temperature

correction from equilibrator temperature to SST is:

fCO2(SST) = fCO2(eq) /

Exp ((Teq-SST) * [0.03107 – 2.7851E-4 * Teq – 1.8391E-3 * ln(fco2eq * 1.0E-6)])

where SST is sea surface temperature and Teq is the equilibrator temperature in

degrees °C.

Sampling Cycle:

The system runs on an hourly cycle during which 3 standard gases, 3 air

samples from the bow tower and 8 surface water samples (from the

equilibrator head space) are analyzed on the following schedule:

Mins. after hour		Sample

4 				Low Standard

8 				Mid Standard

12 				High Standard

16.5				Water

21 				Water

25.5				Water

30				Water

34				Air

38				Air

42				Air

46.5				Water

51				Water

55.5				Water

60				Water

NOTES ON DATA:

Columns have a default value of –999.99 in case of instrument malfunction,

erroneous readings or missing data.  Furthermore, if a suspicious xCO2 value,

pressure or temperature value is encountered, the fCO2 is not calculated.

Analytical Instrument Manufacturer/Model:

The Ron Brown system (version 2.6) was built by Craig Neill in 1999.  The analyzer 	is a LI-COR 6251 (analog output) infrared analyzer.

Standard Gases and Reference Gas:  The three standard gases came from CMDL

in Boulder and are directly traceable to the WMO scale.  While individual data

points above the high standard gas concentration or below the low standard gas

concentration may not be accurate, the general trends should be indicative of the

seawater chemistry.

Description of any additional environmental control:

The system is located in the Hydro Lab of the Ron Brown.  The room is

air-conditioned with little temperature fluctuation.

Resolution of measurement:

The resolution of the instrument is better than 0.1 ppm.

Estimated overall uncertainty of measurement:

The xCO2eq measurements are believed accurate to 0.1 ppm.  The

fCO2@SST measurements are believed to be precise to 0.2 ppm.

List of calibration gases used:

The standards used during the 2007 field season were:

STANDARD		TANK #	CONCENTRATION	VENDOR

Low			CA06745	289.06		ESRL

Mid			CA05398	370.90		ESRL

High			CA05344	411.42		ESRL

High			CC06352	514.29		ESRL

Traceability to an internationally recognized scale (including date/place of last calibration made):

All standards are obtained from NOAA/CMDL, now called the Global

Monitoring Division of the Earth System Research Laboratory and are directly

traceable to WMO scale.

Uncertainty of assigned value of each calibration gas:

The uncertainty based on pre and post cruise calibrations is less

than 0.05 ppm.

Pressure/Temperature/Salinity:

For information about the ship’s thermosalinograph or other sensors, contact Chief 	Survey Tech Jonathan Shannahoff at jonathan.shannahoff@noaa.gov.

Units:

All xCO2 values are reported in parts per million (ppm) and fCO2 values

are reported in microatmospheres (uatm) assuming 100% humidity at the

equilibrator temperature.

Bibliography:

DOE (1994). Handbook of methods for the analysis of the various parameters of the

carbon dioxide system in sea water; version 2. DOE.

Feely, R. A., R. Wanninkhof, H. B. Milburn, C. E. Cosca, M. Stapp and P. P. Murphy

(1998). A new automated underway system for making high precision pCO2

measurements onboard research ships. Analytica Chim. Acta 377: 185-191.

Ho, D. T., R. Wanninkhof, J. Masters, R. A. Feely and C. E. Cosca (1997).

Measurement of underway fCO2 in the Eastern Equatorial Pacific on NOAA

ships BALDRIGE and DISCOVERER, NOAA data report ERL AOML-30, 52 pp.,

NTIS Springfield.

Wanninkhof, R. and K. Thoning (1993). Measurement of fugacity of CO2 in surface

water using continuous and discrete sampling methods. Mar. Chem. 44(2-4):

189-205.

Weiss, R. F. (1970). The solubility of nitrogen, oxygen and argon in water and

seawater. Deep-Sea Research 17: 721-735.

Weiss, R. F. (1974). Carbon dioxide in water and seawater: the solubility of

a non-ideal gas. Mar. Chem. 2: 203-215.

Weiss, R. F., R. A. Jahnke and C. D. Keeling (1982). Seasonal effects of

temperature and salinity on the partial pressure of CO2 in seawater.

Nature 300: 511-513.

Comments related to all 2007 data:

1.	xCO2 values outside the range of the standard gases (i.e. below the low standard or	above the high standard) are not as accurate as values within the range.  However,	the general trends should be indicative of the seawater chemistry.

2.	The standard gases for the first 8 legs (RB200701-RB200708) were 289.06 ppm, 	370.90 ppm, and 411.42 ppm.  The standards for the last 3 legs (RB200707 and	RB200708) were 289.06 ppm, 370.90 ppm, and 514.29 ppm.

Comments related to the individual legs:

RB200701:	1. The water flow to the equilibrator was shut down on April 8 from 				0446 to 1200.  All water values in this time period have been removed.

2. The feeds from the ship's computer system (SCS) and the Micro TSG in 			the Hydro Lab sink did not work and the SCS data was merged later by 				matching maxima and minima in the SST and equilibrator temperature 				data.  The PC clock was found to be ~ 5:45 early at the start of the 				cruise and ~ 7:15 early at the end for a time loss of 1:30 over the 				length of the cruise.  Consequently, 5:45 was added to the PC clock 				time plus an increment equal to the spreadsheet row number divided by 				the total number of rows.  The SCS data was then merged using the 				adjusted PC time.

3. On April 5 and 6 air values were much higher than normal (390 - 420 				ppm). Most of these were likely caused by stack gas contamination but 				not all can be attributed to this source.  Since the ship was far from 				land at this time, all air values during the period have been removed.

RB200702:	1. The water flow to the equilibrator was shut down for 4 hours on 				April 22 from 0300 to 0700.  All water values in this time period have 				been removed.

RB200703:	1. The water flow to the equilibrator was shut down from May 10 at 				23:20 to May 11 at 0700 and from May 21 at 0920 to 1033.  All water 				values in this time period have been removed.

2. The SST readings froze from May 6 at 1842 to May 7 at 1000.  They 				were replaced with the equilibrator temperature - 0.12 degrees C during 			this period.

3. The ship's TSG began giving inaccurate readings on May 19 at 					approximately 1100.  All salinity values for this data set have been 				taken from the Micro TSG in the Hydro lab sink.

4. On May 16 at 1030, the pump that draws air from the bow mast began 				to leak.  All air values after this time have been removed.  When the 				ship came into Ft. Lauderdale, the pump's diaphragm was replaced, 				correcting the problem.

RB200704:	1. Gas flow through the equilibrator dropped to zero on June 8 at 0740 				and never recovered.  All sea surface values after this time have been 				removed.  Air values were not affected.

2. SST and salinity values from the ship's TSG were bad in several 				spots.  Since these are used for computing fCO2 and occurred after the 				flow stopped in the water phase, no attempt has been made to correct 				them.

3. Some atmospheric values above the normal open-ocean value (~ 383 				ppm) have been retained.  There is no indication of stack gas 					contamination or instrument malfunction and as the ship was near land, 				it is assumed that these high values were cause by land-based sources.

RB200705:	1. The ship went into Boston for repairs and the system was shut down 				from July 28 at 1507 GMT to July 29 at 0216 GMT.

2. When coming out of Boston on July 29, the relative wind direction 				sensor was 'stuck' on 140 from 0216 GMT to 2359 GMT.  This made it 				impossible to tell whether high air values were real or caused by stack 			gas.  All high air values in this period have been flagged as 3 					(questionable).

RB200706:	1. The water flow to the equilibrator was shut down from Aug 9 from 				1140 to 1257.  All water values in this time period have been removed.

2. The SST readings froze on Aug. 13 from 1100 to 2047.  They were 				replaced with the equilibrator temperature - 0.245 degrees C during 				this period.

3. The ship's TSG began giving inaccurate readings on Aug. 18 at 					approximately 1200.  All salinity values for this data set have been 				taken from the Micro TSG in the Hydro lab sink.

4. Beginning on Aug. 11 at ~0510 the first of each group of 4 water 				xCO2 values was 0.5 - 2.0 ppm high.  This was apparently caused by gas 				flow that was too high resulting in overpressurization of the Licor 				sample cell.  For the remainder of the cruise ~95% of the first water 				values in each group of 4 measurements has been removed.

RB200707:	No problems of note.

RB200708:	1. The first of each set of 4 water measurements were high due to 				excessive gas flow in the water phase that caused over-pressurization 				of the Licor sample cell.  These values were removed in 454 of 483 				cases.

RB200709:	1. The salinity readings from the ship's TSG were inaccurate and were 				replaced with the salinity readings from the Seabird Micro TSG in the 				Hydro Lab sink.

2. The system was down from Oct. 19 at 1816 through Oct. 23 at 1828 due 			to a computer crash.  When it came back on, water flow was intermittent 			for about 5 hours and many water values in this time period have been 				removed.

3. The ancillary data from the ship's computer system ends on Nov. 6 at 			0010, 30 hours before the end of the cruise.  Since this data includes 				the barometric pressure which is used to compute fCO2Air, the missing 				values have been replaced with EqP + 0.8, where EqP is the pressure 				recorded by the Setra barometer in the Hydro Lab.  The equation is 				based on the difference in readings prior to the end of the ship's 				data.

RB200710:	1. The salinity readings from the ship's TSG were inaccurate and were 				replaced with the salinity readings from the Seabird Micro TSG in the 				Hydro Lab sink.

2. There was no gas flow in the water phase until Nov. 20 at 1448.  All 			values before this time have been removed.

RB200711:	1. The salinity readings from the ship's TSG were inaccurate and were 				replaced with the salinity readings from the Seabird Micro TSG in the 				Hydro Lab sink.

2. The millivolt response of the calibration standards was incorrect 				due to low gas flow which caused incomplete flushing of the Licor 				sample cell.  The problem was corrected on Jan. 9 at 1600 and was 				apparently caused by a gas flow meter that was reading ~40 ml/min too 				high.  During the period from Jan. 2 to Jan. 8, reference gases of 				concentrations 411.42 and 287.95 were passed through the system as 				samples for several hours.  The millivolt response of these gases has 				been used to correct the millivolt response of the standards during the 			period of low gas flow.  All air and equilibrator headspace samples 				prior to Jan. 9 at 1600 have been flagged as 3 or questionable for this 			reason.  Because of the incomplete flushing, gas from the preceding 				measurement phase was present in the sample cell.  For the mid and high 			standards this gas was lower in concentration and caused the voltage 				response to be too low.  For the low standard, the preceding gas was 				equilibrator head space gas and was always higher in concentration, 				although it varied from 345 to 575 ppm over the course of the cruise.  				The low standard voltage response was set to 3.75 mV for the entire 				section of incomplete flushing.  The mid and high responses were 					calculated using the equation adjusted voltage = original voltage + 16 				mV + 0.6mV * (80 - f) where f is the recorded gas flow in ml/min.  The 				last term was added to account for the fact that the higher the gas 				flow (which ranged from 45 to 55 ml/min), the more complete the 					flushing and the less depression in the response.  The corrections were 			arrived at using a process of trial and error and were based on 					responses of recent cruises using the same set of standards.  The xCO2 				values of the standards that were run as samples and the air values 				were then examined to determine the equation that best matched the 				expected values.

3. The following calibration gases were run through the system as 				samples:

Jan. 2	2334 - 2355		411.42 ppm

Jan. 3	1825 - 2100		411.42 ppm

Jan. 6	1111 - 1600		411.42 ppm

Jan. 6	1611 - 2200		287.95 ppm

Jan. 8-9	2211 - 0530		411.42 ppm

4. From Jan. 5 at 1840 to Jan 6 at 1045, gas flow in the water phase 				was low due to an apparent blockage in the tubing from the pump.  The 				first 3 of each set of 4 headspace gas samples has been removed during 				this period.  The data indicates that there was enough gas flow to 				flush the Licor sample cell by the time the fourth sample was measured.

5. Atmospheric air samples were higher than expected (386-405 ppm) for 				about a week after the ship left port and did not fall below 386 ppm 				until around Dec. 23.  As no indication of stack gas contamination can 				be found they are assumed to be caused by shoreside sources and have 				been left in.  As with all samples in this time period they have been 				flagged as 3 (questionable).

6. On Dec. 23 at 1750 the air input line to the bow mast was changed to 			a second air line that had previously been run to see if contamination 				in the line was responsible for the high air values.

7. On Dec. 30 the glass wool in the filter at the air line intake on 				the bow mast was replaced.

8. On Jan. 3 at 2123 the system was shut down and a bad solenoid valve 				on the Licor sample input was replaced.  The system was restarted at 				2201.

9. On Jan 4 at 2314 the water was shut off and the equilibrator 					showerhead was cleaned.  Water was turned back on at 2323.

10. On Jan. 7 at 1528 the system was shut down and the Licor was 					replaced with the spare after first setting its zero and span.  The 				system was turned back on at 2245.

11. On Jan. 8 the system was shut down from 0210 to 2125 to clean the 				rotor of the Valco valve.

12. On Jan. 9 at 1600 the gas flow of the calibration gases was 					increased to ~ 100 ml/min.  This resulted in air values in the expected 			range (~383 ppm).

13. On Jan. 9 from 2045 to 2200 the system was shut down to reinstall 				the original Licor and to replace the nylon tubing to its inputs and 				outputs.